# Manager-Facing Technical Review of the EEG Dashboard

Date: March 30, 2026

Project: EEG / BCI Dashboard Application

Primary codebase: bci_dashboard

## 1. Executive Summary

This application is a working real-time EEG dashboard with live device integration, calibration, metrics, training modules, sessions, and supporting runtime logic. The core system is functional on the current Windows developer setup, and the codebase already contains meaningful optimization work for live rendering, smoothing, buffering, and asynchronous recording.

However, the application is not yet ready for broad rollout across other PCs without setup work. There are confirmed runtime issues, including a training-game crash, repeated physiological calibration failures, and recurring native Bluetooth/device errors in SDK logs. The application also depends on Windows-specific components such as CapsuleClient.dll, os.startfile, and shell-based launch commands, which makes current portability limited.

Manager conclusion: the app is suitable for controlled demos and continued development on the current machine, and it may run on similar Windows PCs with manual setup, but it is not yet packaged or hardened enough for simple deployment to fresh systems or non-technical end users.

## 2. Current Debugging Findings

### Issue 1: Confirmed training gameplay crash

Severity: Critical

Status: Confirmed finding

Root cause:

- In bci_dashboard/gui/training_games.py, the DIR_LABELS map does not include "storm".

- The gameplay logic can set direction = "storm" and later index DIR_LABELS[direction].

Evidence:

- bci_dashboard/gui/training_games.py:23

- bci_dashboard/gui/training_games.py:360

- bci_dashboard/gui/training_games.py:406

- bci_dashboard/logs/dashboard.log contains repeated KeyError: 'storm'

Impact:

- This is a direct live runtime crash during training gameplay.

Manager explanation:

- One game state can produce a label that the UI label map does not know about, which causes the application to throw an exception and break the run.

### Issue 2: Physiological baseline calibration repeatedly times out

Severity: High

Status: Confirmed finding

Root cause:

- The calibration manager runs a staged quick-calibration flow and applies a hard timeout to baseline stages.

- The physiological stage repeatedly times out and then fails after retry.

Evidence:

- bci_dashboard/calibration/calibration_manager.py:67

- bci_dashboard/calibration/calibration_manager.py:103

- bci_dashboard/calibration/calibration_manager.py:453

- bci_dashboard/calibration/calibration_manager.py:460

- bci_dashboard/logs/dashboard.log contains repeated:

- Calibration physiological baseline timed out; retrying

- Calibration physiological baseline timed out after retry

- Calibration failed: Physiological baseline calibration timed out. Please retry.

Impact:

- Quick calibration is unreliable in real runs.

- Training and manager demos can stall before the user reaches stable gameplay.

Manager explanation:

- The app’s final calibration stage does not complete reliably, so users can be blocked even when the rest of the app is working.

### Issue 3: NFB serialization fallback is active because of SDK value-shape problems

Severity: Medium

Status: Confirmed finding

Root cause:

- Some native values coming back from the SDK do not always convert cleanly to integers.

- The calibration path already contains a safety fallback to avoid crashing.

Evidence:

- bci_dashboard/calibration/calibration_manager.py:362

- bci_dashboard/logs/dashboard.log contains repeated warnings like:

- Falling back to safe NFB serialization: invalid literal for int() with base 10: b'...'

Impact:

- The app survives, which is good.

- But the warning indicates a still-unresolved integration mismatch between expected Python values and native SDK values.

Manager explanation:

- The team already added a protective fallback so the application does not crash, but the underlying data-conversion issue is still present.

### Issue 4: Native BLE/device initialization is unstable

Severity: High

Status: Confirmed finding

Root cause:

- The SDK/native layer is repeatedly failing to start scanning or to fully create the BLE device protocol.

Evidence:

- bci_dashboard/logs/Logs/SDK.log contains:

- Failed start BLE scanner. BLE adapter not found or disabled

- Failed read characteristic value: Unreachable

- Failed create Smart Band BLE Protocol: [Can not read Serial number]

- bci_dashboard/logs/Logs/Device.log contains:

- Failed to create BLEDevice: Can not read Serial number

- Execute command 23 failed: Device error: [UNKNOWN]

- Device is not connected!

Impact:

- Device connection and stability are not yet fully reliable.

- Another PC with slightly different Bluetooth conditions may behave worse.

Manager explanation:

- The app’s connection quality depends not only on the code, but also on the Bluetooth environment and the native SDK layer, which currently shows repeated failures in logs.

### Issue 5: Current environment is missing declared dependencies

Severity: High

Status: Confirmed finding

Root cause:

- The repo requirements include pyserial and h5py, but those modules are not currently importable on this machine.

Evidence:

- bci_dashboard/requirements.txt

- Runtime environment check on this machine:

- PySide6: OK

- pyqtgraph: OK

- numpy: OK

- pandas: OK

- serial: FAIL

- h5py: FAIL

- mne: OK

- bci_dashboard/storage/session_recorder.py:152

- bci_dashboard/logs/dashboard.log repeatedly logs:

- h5py is not installed; session.h5 will not be written.

Impact:

- Session recording is partially degraded.

- Prosthetic-arm serial features will not work unless pyserial is installed.

Manager explanation:

- Even on the current machine, the software environment is not fully aligned with the project requirements, which is a warning sign for deployment readiness.

### Issue 6: Core runtime has Windows-only assumptions

Severity: High

Status: Confirmed finding

Root cause:

- The app relies on Windows-native DLL and shell behaviors.

Evidence:

- bci_dashboard/utils/config.py:10 uses CapsuleClient.dll

- bci_dashboard/device/capsule_bridge.py:47 raises if DLL is missing

- bci_dashboard/gui/dashboard_screen.py:779 uses os.startfile

- bci_dashboard/gui/sessions_screen.py:594 uses os.startfile

- bci_dashboard/gui/neuroflow_runtime.py:97 launches via cmd /c start

Impact:

- Current deployment is effectively Windows-only.

- Linux and macOS are not ready for the main application path.

Manager explanation:

- The application can currently be treated as a Windows product only.

### Issue 7: Architecture is tightly coupled and harder to maintain

Severity: Medium

Status: Confirmed finding

Root cause:

- MainWindow owns device orchestration, calibration, routing, session recording, graph history, and page coordination.

- training_games.py is a very large multi-controller module.

Evidence:

- bci_dashboard/gui/main_window.py

- bci_dashboard/gui/training_games.py

Impact:

- Future debugging, onboarding, and feature changes are riskier and slower than they need to be.

Manager explanation:

- The app works, but too much logic lives in a few large modules, which makes long-term stability harder.

### Issue 8: Scan-helper architecture looks stale or partially unused

Severity: Medium

Status: Likely finding / inference

Root cause:

- The codebase still contains a helper-process scanning strategy and QProcess-related cleanup logic, but current scanning appears to use direct locator calls.

Evidence:

- bci_dashboard/device/scan_helper.py

- bci_dashboard/device/device_manager.py

- scan-helper runtime log directory exists but no active helper log files were found

Impact:

- This is more of a maintainability and correctness risk than an immediate crash.

Manager explanation:

- The code shows signs of multiple scanning approaches over time, which suggests cleanup is still needed.

### Issue 9: Legacy/duplicate UI files remain in the repo

Severity: Low

Status: Confirmed finding

Root cause:

- Historical screen variants such as dashboard_screen_old.py and dashboard_screen_new.py are still present.

Evidence:

- Static repo inspection and compile discovery

Impact:

- These are not proven live bugs, but they add confusion and maintenance overhead.

Manager explanation:

- There is some historical code still in the project that should be cleaned up before packaging or handoff.

### Open issues needing runtime verification

Whether the UI-thread callback pump introduces visible lag under heavy live use

Whether repeated calibration-start logging is duplicated user action or duplicate signal flow

Whether reconnect fully restores all classifiers after a real disconnect

Whether physiological timeout is mainly signal-quality related or stage-logic related

## 3. Performance Review

### Current performance behavior

The application is not using a crude busy loop. It is timer- and callback-driven:

Capsule callback pump: `20 ms`

- bci_dashboard/device/capsule_bridge.py:58

Dashboard duration timer: `1000 ms`

- bci_dashboard/gui/dashboard_screen.py:140

Dashboard EEG refresh timer: `50 ms`

- bci_dashboard/gui/dashboard_screen.py:145

Dashboard PSD throttle: `0.08 s`

- bci_dashboard/gui/dashboard_screen.py:126

Training calibration timer: `250 ms`

- bci_dashboard/gui/training_screen.py:317

Training gameplay timer: `250 ms`

- bci_dashboard/gui/training_screen.py:321

Neuroflow simulation timer: `50 ms`

- bci_dashboard/gui/neuroflow_training_page.py

Neuroflow UI refresh timer: `100 ms`

- bci_dashboard/gui/neuroflow_training_page.py

Status/reconnect polling: `5000 ms`

- bci_dashboard/utils/config.py:62

### Likely latency sources

Confirmed findings

Main-thread risks exist because the native Capsule callback pump, PSD snapshot extraction, EEG snapshot extraction, and some page updates are coordinated from the main GUI runtime:

- bci_dashboard/gui/main_window.py:1284

- bci_dashboard/gui/main_window.py:1310

- bci_dashboard/gui/main_window.py:1464

- bci_dashboard/gui/main_window.py:1495

Rendering overhead is partially controlled, but still present:

- the electrode table stores up to 4000 samples per channel

- redraws pyqtgraph plots

- uses downsampling and clip-to-view

- bci_dashboard/gui/widgets/electrode_table.py:38

- bci_dashboard/gui/widgets/electrode_table.py:611

- bci_dashboard/gui/widgets/electrode_table.py:698

Training modules add periodic gameplay work and in some cases audio and richer visuals:

- bci_dashboard/gui/training_screen.py

Likely findings / inferences

If the dashboard, live EEG, PSD processing, and a game are all active together, the single-threaded UI path may become the main source of perceived lag.

The native SDK callback pump being timer-driven on the Qt side is safe and simple, but it may become a bottleneck on slower PCs.

### Logging, buffering, and storage overhead

Confirmed findings

Session writing is already moved off the UI path using a background recorder thread:

- bci_dashboard/gui/main_window.py:79

- bci_dashboard/gui/main_window.py:86

Session data is buffered and flushed in batches:

- bci_dashboard/storage/session_recorder.py

PPG metrics are derived from rolling windows in pure helpers:

- bci_dashboard/gui/raw_metrics.py

### Performance verdict

Real-time suitability: Yes, for controlled Windows use on the current machine

Confidence level: Moderate

Why not higher:

- critical native/device failures are still present

- the most time-sensitive path still shares the GUI thread

- there is no evidence yet from runtime profiling on slower or fresh machines

## 4. Existing Optimizations Identified

### Optimization 1: EMA smoothing in live metrics

Location:

- bci_dashboard/classifiers/emotions_handler.py:25

- bci_dashboard/classifiers/productivity_handler.py:29

What it does:

- smooths incoming emotion and productivity values with exponential moving averages

Why it helps:

- reduces jitter in the dashboard and games

Status:

- useful and active

- slightly increases responsiveness lag, but overall worthwhile

### Optimization 2: Dirty-flag and throttled dashboard updates

Location:

- bci_dashboard/gui/dashboard_screen.py:126

- bci_dashboard/gui/dashboard_screen.py:553

- bci_dashboard/gui/dashboard_screen.py:786

What it does:

- avoids redrawing on every incoming packet

- throttles PSD work and only refreshes when data is marked dirty

Why it helps:

- reduces needless redraw cost

Status:

- effective and important

### Optimization 3: Async session recording

Location:

- bci_dashboard/gui/main_window.py:79

- bci_dashboard/gui/main_window.py:86

What it does:

- records session data through a queue and a background thread

Why it helps:

- prevents recording IO from directly blocking live UI

Status:

- strong optimization already in place

### Optimization 4: Native-packet snapshot extraction before async handoff

Location:

- bci_dashboard/gui/main_window.py:1464

- bci_dashboard/gui/main_window.py:1495

What it does:

- converts native/SDK packet objects into plain Python data before recording or forwarding

Why it helps:

- avoids lifetime/use-after-free issues with native objects

- makes async recording safer

Status:

- important safety and stability optimization

### Optimization 5: Rolling raw-metric aggregation helpers

Location:

- bci_dashboard/gui/raw_metrics.py

What it does:

- centralizes band-history aggregation and PPG metric derivation

Why it helps:

- keeps calculations testable and avoids duplicated UI-side logic

Status:

- good structure improvement already present

### Optimization 6: Reconnect grace period and retry control

Location:

- bci_dashboard/device/device_status_monitor.py

What it does:

- waits for grace period, counts misses, retries on disconnect

Why it helps:

- avoids false disconnect reactions

Status:

- useful, but still limited by native BLE reliability

### Optimization 7: Plot downsampling and bounded EEG buffers

Location:

- bci_dashboard/gui/widgets/electrode_table.py:38

- bci_dashboard/gui/widgets/electrode_table.py:611

- bci_dashboard/gui/widgets/electrode_table.py:698

What it does:

- limits stored samples

- downsamples plots to a smaller point set

- clips rendering to view

Why it helps:

- lowers graph rendering overhead

Status:

- active and useful, but still on the GUI path

### Optimization 8: Calibration safety fallback

Location:

- bci_dashboard/calibration/calibration_manager.py:362

What it does:

- falls back to safer serialization when SDK values behave unexpectedly

Why it helps:

- keeps calibration from crashing

Status:

- good defensive improvement, but underlying issue remains unresolved

## 5. Improvements Already Done

Based on the current code and logs, several meaningful improvements appear to have already been implemented:

safer session recording through async queue/thread handoff

safer native-object handling through snapshot extraction

dirty-flag and throttled dashboard refresh logic

smoothing for noisy real-time metrics

test coverage for dashboard metrics, spectrum widgets, Neuroflow, sessions, and calibration scenarios

defensive calibration serialization fallback

more robust reconnect behavior than a simple binary connected/disconnected flag

graceful degradation when `h5py` is missing

These are real improvements, not cosmetic ones. The codebase is clearly past the earliest prototype stage. The remaining problem is that the app still has a few unresolved failures at exactly the points that matter most for demos and deployment: device reliability, calibration reliability, startup environment consistency, and packaging.

## 6. Future Optimization Plan

### 1. Critical

Fix the confirmed `storm` gameplay crash in `training_games.py`

- Expected benefit: removes a proven live crash

Stabilize physiological baseline calibration with better diagnostics and less opaque timeout handling

- Expected benefit: higher success rate in real sessions

Add startup preflight checks for:

- DLL presence

- SDK importability

- Python dependency availability

- Bluetooth readiness

- writable session/log directories

- Expected benefit: much clearer deployment/startup behavior

Resolve or simplify the BLE scanning architecture

- Expected benefit: fewer device-discovery failures and less support complexity

### 2. High-value

Move more non-UI work off the main thread where safe

- especially PSD-heavy work and expensive snapshot creation

- Expected benefit: better responsiveness under load

Refactor `MainWindow` into smaller coordinating services

- Expected benefit: easier debugging and lower regression risk

Split `training_games.py` into smaller modules by game family

- Expected benefit: easier maintenance and safer changes

Make Neuroflow launch targets configurable instead of hard-coded

- Expected benefit: more reliable app-launch behavior on other PCs

### 3. Medium-term

Build a packaged Windows distribution

- likely PyInstaller or similar

- bundle runtime, assets, DLL, SDK wrappers

- Expected benefit: easier execution on similar Windows PCs

Add log rotation and support-bundle export

- Expected benefit: easier field debugging

Add explicit runtime profiling on:

- dashboard only

- dashboard + training

- dashboard + recording

- Expected benefit: real performance data for managers and QA

### 4. Optional

Investigate true cross-platform support only after Windows packaging is stable

- Expected benefit: future portability

- Current priority: low

## 7. Cross-System Execution Assessment

### Overall answer to the manager question

“Will this application be able to run on other systems easily?”

Current answer: Not easily yet.

It can likely run on another similar Windows PC if the environment is carefully prepared, but it is not currently a plug-and-play product for fresh PCs.

### Can it run on other PCs now?

Yes, under conditions

Windows PC

correct Python version, currently observed as `Python 3.10.9`

required packages installed

local folder layout preserved

`bci_dashboard/lib/CapsuleClient.dll` present

`bci_dashboard/capsule_sdk` present

Bluetooth adapter working and enabled

hardware/device behavior similar to the current setup

No, not easily on fresh systems

no packaged installer

no root deployment guide

current machine itself is missing two declared dependencies

core runtime uses Windows-specific APIs

### Platform assumptions

Confirmed findings

Windows DLL dependency:

- bci_dashboard/utils/config.py:10

Windows file-opening behavior:

- bci_dashboard/gui/dashboard_screen.py:779

- bci_dashboard/gui/sessions_screen.py:594

Windows shell launching:

- bci_dashboard/gui/neuroflow_runtime.py:97

### Python assumptions

Confirmed findings

Current machine uses `Python 3.10.9`

No formal top-level packaging metadata was found for the main app

### Native dependency assumptions

Confirmed findings

`CapsuleClient.dll`

SDK wrapper modules under `bci_dashboard/capsule_sdk`

optional vendored `neurosdk` for prosthetic-arm/BrainBit-related paths

### Hardware and driver assumptions

Confirmed findings

Bluetooth adapter is required for the main headband flow

BLE instability is already visible in native logs

optional Arduino serial connection is required for the prosthetic-arm path

### Folder structure assumptions

Confirmed findings

paths are derived from local app folders

the code repeatedly uses `sys.path.insert(...)` to make SDK and local modules importable

this is workable for development, but weak for deployment

### Cross-platform conclusion

Windows: partially viable with setup

Linux: not ready

macOS: not ready

Cross-platform product claim: not supportable from current code state

## 8. Deployment and Packaging Recommendations

To make the application easier to run on other systems, the following are needed:

### Immediate packaging/setup steps

Pin and document the supported Python version

Ensure `requirements.txt` installs successfully on a clean Windows machine

Add startup checks that fail clearly if:

- CapsuleClient.dll is missing

- capsule_sdk is missing

- pyserial or h5py are missing when needed

- Bluetooth is unavailable

Provide one clear setup document at repo root

Validate startup on a second Windows PC

### Recommended Windows packaging path

Build a Windows executable bundle

Include:

- Python runtime

- bci_dashboard

- capsule_sdk

- lib/CapsuleClient.dll

- assets, music, images

Add a post-start environment validation screen

Create a simple launch shortcut/start script

### For developers/testers

create a reproducible `venv`-based setup

add one command to verify environment health before first launch

keep optional modules clearly marked as optional

## 9. Risks and Blockers

### Technical risks

Critical: confirmed game crash

High: unreliable physiological calibration

High: native BLE/device creation failures

High: undeclared environment drift between requirements and actual installed packages

Medium: main-thread real-time coordination may become a performance bottleneck on weaker systems

Medium: overly large orchestration modules increase change risk

### Architectural risks

tightly coupled runtime ownership in `MainWindow`

stale or duplicated discovery architecture

large training controller file with inconsistent defensive coding

### Platform risks

Windows-specific runtime path

native DLL dependency

Bluetooth/driver sensitivity

optional serial and vendor-library dependencies

### Operational risks

another PC may fail before user reaches dashboard due to environment mismatch

native logs already show hardware/driver-level instability

without packaging, deployment support burden stays high

## 10. Final Verdict

### Is the application ready for broader usage?

No, not yet.

### Is it ready only for the current machine?

Mostly yes. The current machine is the highest-confidence environment because it already contains the project layout, SDK files, and most of the working runtime context.

### Is it ready for similar Windows systems?

Partially, with manual setup and validation. It is not yet “easy to execute” in the manager’s sense.

### What must be fixed before rollout?

fix the confirmed gameplay crash

improve calibration reliability

harden BLE/device startup and diagnostics

align environment dependencies with actual runtime needs

package the Windows runtime properly

document startup/setup clearly

## Manager-Friendly Summary Paragraph

The EEG dashboard is already a real working application with live EEG integration, calibration, dashboards, games, and session recording, and it includes several real optimizations that make it usable on the current development system. But it is still not ready for simple rollout to other PCs because it depends on a Windows-specific native SDK setup, shows recurring Bluetooth/device and calibration failures in logs, and still has at least one confirmed runtime crash in gameplay. For now, it is best treated as a controlled Windows demo/development application rather than a finished portable product.

## Technical Engineering Summary Paragraph

The current runtime architecture is a Qt-driven, callback-and-timer-based BCI application centered on MainWindow, with native Capsule integration, classifier wrappers, throttled dashboard rendering, and asynchronous session recording. Existing optimizations are meaningful and include EMA smoothing, dirty-flag redraw control, pyqtgraph downsampling, async recording, and safer native snapshot extraction. The main engineering gaps are runtime stability at calibration/device boundaries, UI-thread ownership of key real-time paths, monolithic orchestration, inconsistent defensive handling across controllers, and lack of a reproducible packaging/deployment layer.

## Short Action Plan

Fix the `storm` gameplay crash immediately.

Add deeper diagnostics around physiological calibration timeout and reproduce it with a fresh live session.

Add startup preflight checks for DLLs, Python dependencies, Bluetooth, and writable directories.

Build and validate a packaged Windows version on a second PC before any wider rollout.